AccuSync

Software Requirements

Software Requirements is a living document; please note major changes to this document in the table below.

| Rev. | Author | Jama Baseline Ref | Change order number/Changes |
| --- | --- | --- | --- |
| 01 | Haidee Kachniewicz | B1 - 01/23/2026 | DCO#73885 Initial Release |

Table of Contents

1.	Purpose	4

2.	Scope	4

3.	References	4

4.	Definition References	4

5.	Software Requirements	5

5.1	General	5

5.2	Login	6

5.3	Patient Management	7

5.4	Import	9

5.5	Export	10

5.6	OAE Test Result	12

5.7	ABR Test Result	12

5.8	User Account Management	12

5.9	Profile Management	13

5.10	Device Management	14

5.11	Site Management	15

5.12	Facility Management	16

5.13	Location Management	17

5.14	ABR Test Protocol Configuration	17

5.15	DPOAE Test Protocol Configuration	18

5.16	Risk Factor Configuration	18

5.17	Comments Configuration	19

5.18	Patient Field Configuration	19

5.19	Report Generation	20

5.20	Language	20

5.21	S4H	22

5.22	AccuScreen Pro Device Communication	24

5.23	Network	25

5.24	Operating System	25

5.25	Alarms, Warnings, Operator Messages	25

5.26	Audit Trail	26

5.27	Logging	27

5.28	About	28

5.29	Help	28

5.30	Database Storage	28

5.31	Firmware Update	29

5.32	Installation	29

# Purpose

The purpose of this document is to define the software requirements for the AccuSync Data Management Software in accordance with the user requirements and the quality management system.

# Scope

The scope of this software requirements document covers all of the functionalities of the AccuSync Data Management Software. The software will integrate with the AccuScreen Pro device and other systems, providing interfaces for data import, export, and configuration operations.

# References

| Document number | Description |
| --- | --- |
| QMS-000075_SENSORY | Corporate Design & Development Policy |
| QMS-003712_SENSORY | Design Responsibilities & Definitions |
| DOC-075955 | AccuScreen Pro User Requirements |
| DOC-076750 | AccuScreen Pro Product Requirements |
| DOC-076988 | AccuScreen Pro Risk Management Plan |
| DOC-076518 | AccuScreen Pro Risk Assessment Spreadsheet |
| TBD | AccuScreen Pro Usability Engineering File |
| DOC-077042 | AccuScreen Pro Regulatory Plan |

# Definition References

Category Definition:

| H | Hardware (Mechanical and/or electronic) |
| --- | --- |
| F | Firmware |
| S | Software |
| U | User interface (software or hardware) not requiring risk mitigation. |
| U* | User interface requiring mitigating a risk (add HZ ID) and must be listed in the Usability file. |
| I |  |
| E | Essential Design Output (EDO) |
| R | Regulatory / Standard |

Hazard definition (HZ ID):

If the requirement is not mitigating a risk, it should show “None”.

| Definitions | Description |
| --- | --- |
| ABR | Auditory Brainstem Response |
| CSV | Comma-Separated Values |
| DPOAE | Distortion Product Otoacoustic Emissions |
| JSON | JavaScript Object Notation |
| NHSP | Newborn Hearing Screening Programme |
| OAE | Otoacoustic Emissions |
| S4H | SMaRT4Hearing |
| TEOAE | Transient Evoked Otoacoustic Emissions |
| USB | Universal Serial Bus |
| XML | Extensible Markup Language |

# Software Requirements

## General

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254873 | None | S,U | Data Management Options | The software shall provide data management options Add, Edit, and Delete for creating, modifying, and removing database records. |
| GID-254874 | None | S,U | Add Function | The software shall allow the user to create a new record by selecting Add, which shall open a blank form for data entry. |
| GID-254875 | None | S,U | Edit Function | The software shall allow the user to modify a selected record by selecting Edit, which shall open the record in an editable form. |
| GID-254876 | None | S,U | Delete Function | The software shall allow the user to delete a selected record by selecting Delete, and shall require user confirmation before removal from the database. |
| GID-254877 | None | S,U | Data Management Screen Availability | The software shall provide data management options (Add, Edit, Delete) on the following screens: Patient, User, Profile, Device, Site, Facility, Location, ABR Protocols, DPOAE Protocols, Risk Factors, Comments. |
| GID-254878 | None | S,U | Unsaved Data Management Options | The software shall provide unsaved data management options Save, Revert, and Undo for managing uncommitted changes. |
| GID-254879 | None | S,U | Save Function | The software shall allow the user to save all modified data by selecting Save, which shall persist changes to the database. |
| GID-254880 | None | S,U | Revert Function | The software shall allow the user to discard all unsaved changes by selecting Revert, restoring the data to its last saved state. |
| GID-254881 | None | S,U | Undo Function | The software shall allow the user to reverse the most recent unsaved change by selecting Undo. |
| GID-254882 | None | S,U | Unsaved Data Management Screen Availability | The software shall provide unsaved data management options (Save, Revert, Undo) on the following screens: Patient, User, Profile, Device, Site, Facility, Location, ABR Protocols, DPOAE Protocols, Risk Factors, Comments, Patient Field Configuration. |

## Login

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255015 | None | S,U | User Login | The software shall require users to enter username and password credentials to login to the system. |
| GID-255016 | None | S,U | User Logout | The software shall provide an option to logout of the system. |
| GID-255017 | None | S,U | User Account Lockout | The software shall lock out the user from logging into the system after a minimum of 5 failed sequential login attempts with improper credentials. |
| GID-255018 | TBD | S,U* | Credential Encryption | The software shall securely hash all login passwords and encrypt all usernames stored in the system. |
| GID-255019 | None | S,U | Role-Based Access Control | The software shall support role-based access control with minimum two roles: Admin and Screener. |
| GID-255020 | None | S,U | Password Complexity Requirements | The software shall enforce password complexity requirements each user account: • At least 8 characters • At least one upper case letter (A to Z) • At least one lower case letter (a to z) • At least one numerical digit • The password cannot be the same as the three previously used passwords |
| GID-255021 | None | S,U | Deactivated User Login Prevention | The software shall prevent deactivated users from logging in. |
| GID-255022 | None | S,U | Deactivated User Transfer Prevention | The software shall prevent transfer of deactivated user accounts to the device. |

## Patient Management

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254883 | None | S,U | Create Patient Record | The software shall allow users to create a patient record containing patient demographics, caregiver data, medical data, and consent information. |
| GID-254884 | None | S,U | View Patient List | The software shall allow users to view a list of patient records saved in the system. |
| GID-254885 | None | S,U | Edit Patient Record | The software shall allow users to select a patient from the list view and edit the patient record, unless locked by permissions. |
| GID-254886 | None | S,U | Delete Patient Record | The software shall allow administrative users to delete a patient record. |
| GID-254887 | None | S,U | Patient Search | The software shall provide search functionality on the patient list view, allowing users to search by patient ID, first name, last name, date of birth, or date of test range, and display the filtered results. |
| GID-254888 | None | S,U | Patient Risk Factor Selection | The software shall provide an option to set risk factors in the patient details view with the following values: • Yes • No • Unknown |
| GID-254889 | None | S,U | Patient Comment Assignment | The software shall provide an option to assign predefined or patient-specific comments to a patient record. |
| GID-254890 | None | S,U | Patient List Display Fields | The software shall display the following information for each patient in the patient list view: • Patient ID / Hospital ID • Last Name • First Name • Date of Birth • Risk • Comment |
| GID-254891 | None | S,U | View Patient Information | The software shall allow users select a patient from the list view and view patient information. |
| GID-254892 | None | S,U | Mandatory Field Indication | The software shall indicate which fields are mandatory during patient data entry. |
| GID-254893 | None | S,U | Patient Save Validation | The software shall allow the user to save patient information when all mandatory fields are populated. |
| GID-254894 | None | S,U | Patient Test Report Generation | The software shall generate and print predefined patient test reports, including selected tests or all tests. |
| GID-254895 | None | S,U | Delete Patient Test Entry | The software shall allow administrative users to delete individual test entries from a patient's test result list. |
| GID-254896 | None | S,U | Test Result Reassignment | The software shall allow administrative users to reassign test results to the correct patient in the system. |
| GID-254897 | None | S,U | Patient Test List Display Fields | The software shall display the following information for each test of a selected patient: • Test Type • Left Ear Result • Right Ear Result • Date/Time of Test • Test Configuration • Duration • Examiner |

## Import

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254900 | None | S,U | Supported Import Formats | The software shall support the following formats to import: • AccuSync XML • AccuSync JSON • ALGO 5 XML • ALGO Pro JSON • AccuLink XML [Additional formats as required] |

## Export

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254898 | None | S,U | Supported Export Formats | The software shall support the following export formats: • AccuSync JSON • AccuSync XML • HiTrack • OZ • CSV • ALGO 5 XML [Additional formats as required] |
| GID-254899 | None | S,U | Patient Data Export | The software shall export patient and test data to supported external systems via file or web service. |
| GID-256274 | None | S,U | Configurable Export Location | The software shall allow users to specify the folder location for data export. |
| GID-256275 | None | S,U | Export New Patients with Data | The software shall allow export of new patients with data. |
| GID-256276 | None | S,U | Export All Patients | The software shall allow export of all patients. |
| GID-256277 | None | S,U | Export Selected Entries | The device shall allow export of individually selected entries. |
| GID-256278 | None | S,U | Export by Test Date Range | The software shall allow export of entries within a specified test date range. |
| GID-256279 | None | S,U | De-identified Data Export | The software shall provide an export option to remove patient demographics and identifiers from the test data prior to sharing the data or for service review. |
| GID-256280 | None | S,U | AccuSync XML File Storage and Naming Convention | The software shall store AccuSync XML files in ExportData/XML/ folder and named AccuSync_YYYY_MM_DD_HH_MM_SS.xml. The date and time should indicate when the AccuSync XML file was created. |
| GID-256281 | None | S,U | AccuSync JSON File Storage and Naming Convention | The software shall store AccuSync JSON files in ExportData/JSON/ folder and named AccuSync_YYYY_MM_DD_HH_MM_SS.json. The date and time should indicate when the AccuSync JSON file was created. |
| GID-256282 | None | S,U | HiTrack File Storage and Naming Convention | The software shall store HiTrack files for the state in ExportData/SummaryFiles/YYYY/YYYY-MON/ folder and named state_yyyymmdd_hhmmss.txt. A duplicate file named INTHS.txt shall be stored here as well as in the folder HiTrackExportData. The date and time should indicate when the HiTrack file was created. MON indicates a three letter month abbreviation. |
| GID-256283 | None | S,U | OZ Export File Storage and Naming Convention | The software shall store OZ fixed width files for the state in ExportData/OZ/ folder and named state_yyyymmdd_hhmmss.txt. The date and time should indicate when the OZ file was created. |
| GID-256284 | None | S,U | CSV Export File Storage and Naming Convention | The software shall store CSV fixed width files in ExportData/CSV/ folder and named CSV_yyyymmdd_hhmmss.csv. The date and time should indicate when the CSV file was created. |
| GID-256285 | None | S,U | ALGO 5 XML Export File Storage and Naming Convention | The software shall store ALGO 5 XML files in ExportData/XML/ folder and named ALGO5_YYYY_MM_DD_HH_MM_SS.xml. The date and time should indicate when the ALGO 5 XML file was created. |
| GID-256286 | None | S,U | Configurable Export Format | The software shall allow users to select a default export format. |
| GID-256512 | None | S,U | Default Export Format | The default export format shall be AccuSync JSON. |

## OAE Test Result

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254901 | None | S,U | TEOAE Test Result | The software shall allow users to view patient test result details, waveform data, test comments, and device information for the TEOAE test. |
| GID-254902 | None | S,U | DPOAE Test Result | The software shall allow users to view patient test result details, waveform data, test comments, and device information for the DPOAE test. |

## ABR Test Result

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254903 | None | S,U | ABR Test Result | The software shall allow users to view patient test result details, waveform data, test comments, and device information for the ABR test. |

## User Account Management

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254904 | None | S,U | View User List | The software shall allow administrative users to view a list of user accounts. |
| GID-254905 | None | S,U | Create User Account | The software shall allow administrative users to create a user account. |
| GID-254906 | None | S,U | Edit User Account | The software shall allow administrative users to edit a user account. |
| GID-254907 | None | S,U | Unlock User Account | The software shall allow administrative users to unlock a user account. |
| GID-254908 | None | S,U | Delete User Account | The software shall allow administrative users to delete a user account. |
| GID-254909 | None | S,U | User Account Activation | The software shall allow administrative users to activate/deactivate a user account. |
| GID-254910 | None | S,U | User Language Selection | The software shall allow administrative users to select the display language per user. |
| GID-254911 | None | S,U | Account Lockout Duration Configuration | The software shall allow administrative users to configure the lockout duration after 5 consecutive lockouts. |
| GID-254912 | None | S,U | Password Complexity Configuration | The software shall allow administrative users to configure the password complexity (None, Simple, Complex). |

## Profile Management

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254913 | None | S,U | View Profile List | The software shall allow administrative users to view a list of profiles and feature access rights. |
| GID-254914 | None | S,U | Create Profile | The software shall allow administrative users to create a profile. |
| GID-254915 | None | S,U | Edit Profile | The software shall allow administrative users to edit a profile. |
| GID-254916 | None | S,U | Delete Profile | The software shall allow administrative users to delete a profile. |
| GID-254917 | None | S,U | Profile Display Fields | The software shall display the following information for a selected profile: • Name • Description • Components and Permissions |
| GID-254918 | None | S,U | Profile Permission Configuration | The software shall allow administrative users to configure permissions for each profile. |

## Device Management

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254919 | None | S,U | Device List Display Fields | The software shall allow administrative users to view a list of devices with the following information: • Name • Serial number • Last Seen |
| GID-254920 | None | S,U | Add Device | The software shall allow administrative users to add a new device. |
| GID-254921 | None | S,U | Edit Device | The software shall allow administrative users to edit device information. |
| GID-254922 | None | S,U | Delete Device | The software shall allow administrative users to delete a device. |
| GID-254923 | None | S,U | Device User Assignment | The software shall allow administrative users to assign users to each device. |
| GID-254924 | None | S,U | Device Facility Assignment | The software shall allow administrative users to assign facilities to each device. |
| GID-254925 | 6.4 | S,U* | Device Settings Configuration | The software shall allow administrative users to configure the following device settings: • Display timeout • Power timeout • Calibration/pause time • Result terminology • Automatic deletion • ABR Autostart • TEOAE Probe Fit Assistant |
| GID-254926 | None | S,U | Device System Information Display | The software shall display the following System Information for a selected device: • Last Seen • Last Updated • Hardware Version • Firmware Version |

## Site Management

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254927 | None | S,U | View Site List | The software shall allow administrative users to view a list of sites and associated details. |
| GID-254928 | None | S,U | Site Detail Fields | The software shall allow administrative users to enter the following site details: • Name • Description • Code |
| GID-254929 | None | S,U | Add Site | The software shall allow administrative users to add a new site. |
| GID-254930 | None | S,U | Edit Site | The software shall allow administrative users to edit site details. |
| GID-254931 | None | S,U | Delete Site | The software shall allow administrative users to delete a site. |

## Facility Management

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254932 | None | S,U | View Facility List | The software shall allow administrative users to view a list of facilities and associated details. |
| GID-254933 | None | S,U | Facility Detail Fields | The software shall allow administrative users to enter the following facility details: • Name • Description • Code • Site • Location Type |
| GID-254934 | None | S,U | Add Facility | The software shall allow administrative users to add a facility. |
| GID-254935 | None | S,U | Edit Facility | The software shall allow administrative users to edit a facility. |
| GID-254936 | None | S,U | Delete Facility | The software shall allow administrative users to delete a facility. |

## Location Management

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254937 | None | S,U | View Location List | The software shall allow administrative users view a list of locations and associated details |
| GID-254938 | None | S,U | Location Detail Fields | The software shall allow administrative users to enter the following location details: • Name • Description • Code |
| GID-254939 | None | S,U | Add Location | The software shall allow administrative users to add a new location. |
| GID-254940 | None | S,U | Edit Location | The software shall allow administrative users to edit a location. |
| GID-254941 | None | S,U | Delete Location | The software shall allow administrative users to delete a location. |
| GID-255465 | None | S,U | Assign Location | When adding a new location, the software shall provide an option to assign the location to all facilities. |

## ABR Test Protocol Configuration

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254942 | None | S,U | Create ABR Protocol | The software shall allow administrative users create ABR test protocol configurations. |
| GID-254943 | None | S,U | View ABR Protocol List | The software shall allow administrative users view a list of ABR test protocol configurations. |
| GID-254944 | None | S,U | Edit ABR Protocol | The software shall allow administrative users to edit an existing ABR test protocol configuration. |
| GID-254945 | None | S,U | Delete ABR Protocol | The software shall allow administrative users to delete an existing ABR test protocol configuration. |

## DPOAE Test Protocol Configuration

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254946 | None | S,U | Create DPOAE Protocol | The software shall allow administrative users create DPOAE test protocol configurations. |
| GID-254947 | None | S,U | View DPOAE Protocol List | The software shall allow administrative users view a list of DPOAE test protocol configurations. |
| GID-254948 | None | S,U | Edit DPOAE Protocol | The software shall allow administrative users to edit an existing DPOAE test protocol configuration. |
| GID-254949 | None | S,U | Delete DPOAE Protocol | The software shall allow administrative users to delete an existing DPOAE test protocol configuration. |

## Risk Factor Configuration

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254950 | None | S,U | Create Risk Factor | The software shall allow administrative users to create predefined risk factors that can be assigned to patient records. |
| GID-254951 | None | S,U | View Risk Factor List | The software shall allow administrative users to view a list of predefined risk factors that can be assigned to patient records. |
| GID-254952 | None | S,U | Edit Risk Factor | The software shall allow administrative users to edit predefined risk factors that has not been assigned to a patient record. |
| GID-254953 | None | S,U | Delete Risk Factor | The software shall allow administrative users to delete a list of predefined risk factors that is not assigned to a patient record. |
| GID-254954 | None | S,U | Risk Factor Translation | The software shall allow administrative users to enter translated text for risk factor names and descriptions in each supported language. |

## Comments Configuration

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254955 | None | S,U | Create Comment | The software shall allow administrative users to create a comment that can be assigned to a patient record. |
| GID-254956 | None | S,U | View Comment List | The software shall allow administrative users to view a list of predefined comments that can be assigned to a patient record. |
| GID-254957 | None | S,U | Edit Comment | The software shall allow administrative users to edit a predefined comment that can be assigned to a patient record. |
| GID-254958 | None | S,U | Delete Comment | The software shall allow administrative users to delete a predefined comment that is not assigned to a patient record. |
| GID-254959 | None | S,U | Comment Translation | The software shall allow administrative users to enter translated text for comments in each supported language. |

## Patient Field Configuration

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254960 | None | S,U | Patient ID Format Validation Configuration | The software shall allow administrative users to configure patient ID rule format validation. |
| GID-254961 | None | S,U | Patient Mandatory Field Configuration | The software shall allow administrative users to configure which patient record fields are mandatory during data entry. |
| GID-254962 | None | S,U | Patient Active Field Configuration | The software shall allow administrative users to configure which patient record fields are active during data entry. |
| GID-254963 | None | S,U | Patient List Sort Configuration | The software shall allow administrative users to define how patient lists are sorted. |
| GID-254964 | None | S,U | Data Confirmation | The software shall allow configuration of the following user notifications: • Confirmation of saving • Confirmation of deletion • Data modification warning |
| GID-254965 | None | S,U | Report Configuration | The software shall provide an option to add logo to the layout of the report by using a graphics file and to select paper format. |
| GID-254966 | None | S,U | Data Validation | The software shall validate user inputs and prevent saving of invalid or incomplete data according to configured validation rules. |
| GID-255466 | None | S,U | Custom Field Configuration | The software shall provide customizable fields labeled "Available Field" that allow users to rename the fields. |

## Report Generation

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255112 | None | S,U | Patient Report Generation | The software shall allow users to generate patient reports that include demographic information and test result details. |
| GID-255113 | None | S,U | Export Reports | The software shall allow exporting printed reports to supported file formats. |

## Language

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254967 | None | S,U | Default Language Selection | The software shall allow the user to select the default language from a list of supported languages. |
| GID-254968 | None | S,U | Supported Languages | The software shall display the following languages to select as the default language: Phase 1 • English • French • Italian • German • Spanish Phase 2 • Brazilian Portuguese • Chinese (simplified) • Chinese (traditional) • Japanese Phase 3 • Norwegian • Danish • Finnish • Swedish • Turkish [Additional languages as required] |
| GID-256513 | None | S,U | Default Language | The default language setting shall be English. |
| GID-254969 | None | S,U | Localized Display | The software shall display all on-screen text, messages, menus, prompts, and results in the configured language. |
| GID-254970 | None | S,U | Language Change Confirmation | The software shall require user confirmation before applying the selected language as the default language for the application. |

## S4H

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254971 | None | S,U | UK Terminology Standards | The device shall use English (UK) terminology as defined by NHSP, including: Pass = Clear Response (CR), Refer = No Clear Response (NCR), Surname = Last Name, and Forename = First Name. |
| GID-254972 | None | S,U | Synchronize Identifiers and Facility Lists | The software shall synchronize the following fields with the device: • Site Identifier • Device Identifier • List of inpatient facilities • List of outpatient facilities |
| GID-254973 | None | S,U | Site identifier Configuration | The software shall provide an option to configure the Site identifier. |
| GID-254974 | None | S,U | SEDQ Web Service URL | The software shall provide an option to configure the SEDQ Web Service URL . |
| GID-254975 | None | S,U | User List Synchronization | The software shall receive the user list from the SEDQ web service. |
| GID-254976 | None | S,U | Test Results Transfer | The software shall support transfer of screening test results, including binary waveform data, into S4H. |
| GID-254977 | None | S,U | Restrict User Configration | The software shall prevent administrative users from adding a user account. |
| GID-254978 | None | S,U | Default Profile Assignment | The software shall provide an option to select a default user profile to assign to all new users imported from S4H. |
| GID-254979 | None | S,U | Default Password Assignment | The software shall provide an option to set the default password to apply to all new users dowloaded from S4H. |
| GID-254980 | None | S,U | Automatic User Deactivation | The software shall automatically set the account status to "Inactive" for any active user accounts whose identifiers are not present in the received synchronisation data. |
| GID-254981 | None | S,U | Automatic User Reactivation | The software shall automatically restore the account status to "Active" for any inactive user accounts whose identifiers are present in the received synchronisation data. |
| GID-254982 | None | S,U | User Attribute Synchronisation | The software shall retrieve and store Username and User ID from the synchronisation data. |
| GID-254983 | None | S,U | Read-Only User Attributes | The software shall prevent modification of Username and User ID except through synchronisation updates. |
| GID-254984 | None | S,U | Automatic Device Deactivation | The software shall retrieve and store Device ID, Name, and Serial Number from the synchronisation data. |
| GID-254985 | None | S,U | Automatic Device Reactivation | The software shall prevent modification of Device ID, Name, and Serial Number except through synchronisation updates. |
| GID-254986 | None | S,U | Device Attribute Synchronisation | The software shall automatically set the device status to "Inactive" for any active devices whose identifiers are not present in the received synchronisation data. |
| GID-254987 | None | S,U | Read-Only Device Attributes | The software shall automatically restore the device status to "Active" for any inactive devices whose identifiers are present in the received synchronisation data. |
| GID-254988 | None | S,U | Automatic Facility Deactivation | The software shall automatically set the facility status to "Inactive" for any active facilities whose identifiers are not present in the received synchronisation data. |
| GID-254989 | None | S,U | Automatic Facility Reactivation | The software shall automatically restore the facility status to "Active" for any inactive facilities whose identifiers are present in the received synchronisation data. |
| GID-254990 | None | S,U | Facility Attribute Synchronisation | The software shall retrieve and store Facility ID, Name, and Type from the synchronisation data. |
| GID-254991 | None | S,U | Read-Only Facility Attributes | The software shall prevent modification of Facility ID, Name, and Type except through synchronisation updates. |
| GID-254992 | None | S,U | Automatic Risk Factor Deactivation | The software shall automatically set the risk factor status to "Inactive" for any active risk factors whose identifiers are not present in the received synchronisation data. |
| GID-254993 | None | S,U | Automatic Risk Factor Reactivation | The software shall automatically restore the risk factor status to "Active" for any inactive risk factors whose identifiers are present in the received synchronisation data. |
| GID-254994 | None | S,U | Risk Factor Attribute Synchronisation | The software shall retrieve and store Risk Factor ID and Risk Factor Value from the synchronisation data. |
| GID-254995 | None | S,U | Read-Only Risk Factor Attributes | The software shall prevent modification of Risk Factor ID and Risk Factor Value except through synchronisation updates. |

## AccuScreen Pro Device Communication

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-254996 | None | S,F,U | Device Connection | The software shall establish connections with AccuScreen Pro devices through USB. |
| GID-254997 | None | S,F,U | Connection Status | The software shall display connection status indicating whether devices are connected, disconnected, or in the process of connecting. |
| GID-254998 | None | S,F,U | Device Status | The software shall display current status information from connected AccuScreen Pro devices including firmware version and device identification. |
| GID-254999 | None | S,F,U | Patient Record Transfer to Device | The software shall transfer patient demographic information to a connected AccuScreen Pro device. |
| GID-255000 | None | S,F,U | Test Result Transfer from Device | The software shall import patient demographic information and test results from a connected AccuScreen Pro device. |
| GID-255001 | None | S,F,U | Site Configuration Transfer | The software provide an option to transfer Site configuation data to a connected AccuScreen Pro device. |
| GID-255002 | None | S,F,U | Facility Configuration Transfer | The software provide an option to transfer Facility configuration data to a connected AccuScreen Pro device. |
| GID-255003 | None | S,F,U | ABR Protocol Configuration Transfer | The software shall transfer ABR test protocol configurations to a connected AccuScreen Pro device. |
| GID-255004 | None | S,F,U | DPOAE Protocol Configuration Transfer | The software shall transfer DPOAE test protocol configurations to a connected AccuScreen Pro device. |

## Network

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255006 | None | S,U | Network Requirement Check | The software shall require network access when exchanging data with web-service-based systems. |

## Operating System

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255007 | None | S,U | Operating System Compatibility | The software shall operate on Windows 11 (Pro and Enterprise editions). |

## Alarms, Warnings, Operator Messages

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255008 | None | S,U | Authentication Warnings | The software shall display warning messages after repeated failed login attempts and notify users when accounts are locked. |
| GID-255009 | None | S,U | Unsaved Data Warning | The software shall warn users before closing windows or navigating away when unsaved data would be lost. |
| GID-255010 | None | S,U | Connection Status Alerts | The software shall provide visual indicators showing whether a device is connected or disconnected. |
| GID-255011 | None | S,U | Firmware Update Warnings | The software shall display warnings and precautions before initiating device firmware updates. |
| GID-255012 | None | S,U | Validation Error Messages | The software shall display error messages identifying which fields contain invalid data and what corrections are needed. |
| GID-255013 | None | S,U | Operation Status Messages | The software shall display progress indicators and status messages during time-consuming operations such as printing, data transfer, and device firmware updates. |
| GID-255014 | None | S,U | Session Timeout Warning | The software shall display a warning message before automatically logging out users due to inactivity. |

## Audit Trail

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255023 | None | S | User Activity Audit Logging | The software shall maintain an audit log of all user login/logout, login attempts, failures, and password changes. |
| GID-255024 | None | S | Patient Record Audit Logging | The software shall maintain an audit log of all patient record creation, modification, deletion, export, and import events. |
| GID-255025 | None | S | Configuration Change Audit Logging | The software shall shall maintain an audit log of all configuration and setting changes. |
| GID-255026 | None | S | Firmware Upgrade Audit Logging | The software shall maintain an audit log of all firmware upgrades. |
| GID-255027 | None | S | Audit Log Entry Fields | The software shall associate all audit log entries with user ID, timestamp, description, device serial number or identifier, and status of the operation. |
| GID-255028 | None | S | Synchronisation Audit Logging | The software shall maintain an audit log of all successful and failed synchronisation attempts. |
| GID-255029 | None | S | Audit Log Retention | The software shall retain audit logs for a minimum of one year. |
| GID-255030 | None | S | User Status Audit Logging | The software shall generate an audit trail entry for each automatic user status change, including user identifier, timestamp, and triggering synchronisation event. |

## Logging

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255031 | None | S | Log Message Generation | The software shall provide a mechanism to generate log messages during the operation of the software. |
| GID-255032 | None | S | Log Message Types | The software shall provide a logging mechanism that supports informational, warning, and error message types. |
| GID-255033 | None | S | Log Message Storage | The software shall store generated log messages on the system for troubleshooting purposes. |
| GID-255034 | None | S | Log Data Privacy | The software shall not include or retain any patient-identifiable information in the log data. |
| GID-255035 | None | S | Log File Retention | The software shall retain log files for a minimum of one year. |
| GID-255036 | None | S | Log File Security | The software shall prevent unauthorized access to log files. |
| GID-255037 | None | S | Log Protection | The software shall prevent deletion of log files by any user. |
| GID-255038 | None | S | Log Entry Fields | The software shall associate all log entries with timestamp, description and status of the operation. |

## About

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255039 | None | S,U | About | The software shall display information about the application, including the version number, manufacturer name and website. |

## Help

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255040 | None | S,U | Help Documentation | The software shall provide access to user help documentation from within the application. |
| GID-255041 | None | S,U | Context Help | The software shall provide context-sensitive help accessible via the Help icon. |

## Database Storage

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255042 | None | S | Patient Record Storage | The software shall store and maintain patient test records in a local patient database on the system. |
| GID-255043 | None | S | Device Configuration Storage | The software shall store and maintain device configurations in a local settings database on the system. |
| GID-255044 | None | S | AccuScreen Pro Patient Record Storage | The software shall store and maintain patient test records received from the AccuScreen Pro device in a local patient database. |
| GID-255045 | None | S | AccuScreen Pro Device Configuration Storage | The software shall store and update device configurations received from the AccuScreen Pro device in a local setting database. |
| GID-256510 | None | S | Patient Database Encryption | The software shall encrypt all patient test records stored in the local patient database (minimum AES-256). |
| GID-256511 | None | S | Settings Database Encryption | The software shall encrypt all settings configurations stored in the local settings database (minimum AES-256). |

## Firmware Update

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255005 | None | S,F,U | Firmware Update | The software shall be able to perform a firmware update to a connected AccuScreen device. |

## Installation

| PR ID | Hazard ID | Category | Title | Requirement |
| --- | --- | --- | --- | --- |
| GID-255046 | None | S,U | Installation | The software shall provide a guided installation process that verifies prerequisites and system requirements before installation. |
| GID-255047 | None | S,U | Uninstallation | The software shall provide an uninstall process that removes all application components with an option to preserve or delete database files. |